Commissioning of a Novel MicroCT/RT System for Small Animal Conformal Radiotherapy

Manuel Rodriguez1,2, Hu Zhou1, Paul Keall1,  Edward Graves1

1Department of Radiation Oncology, Stanford University, CA, USA, 2 Departamento de Fisica, Universidad Nacional Autonoma de Honduras, Honduras.

Contact address:

Manuel Rodriguez

Stanford University Cancer Center

875 Blake Wilbur Drive, G200

Stanford, CA 94305

Phone (650) 721-5475

Email: mrrodriguezv@gmail.com

#### Conflict of Interest Notification

The authors declare no conflicts of interest related to this work.

ABSTRACT

The purpose of this work was to commission a 120 kVp photon beam produced by a micro-computed tomography (microCT) scanner for use in irradiating mice to therapeutic doses. A variable-aperture collimator has been integrated with the system to allow the delivery of beams with pseudocircular profiles of arbitrary width between 0.1 and 6.0 cm. The dose rate at the isocenter of the system was measured using ion chamber and gafchromic EBT film as 2.13-1.56 Gy/min at the surface of water for field diameters between 6.0 and 0.2 cm. The dose rate decreases approximately 10% per every 5 mm depth in water for field diameters between 1.0 and 0.5 cm.  The flatness, symmetry and penumbra of the beam are 3.6%, 1.0% and 0.5 mm respectively. These parameters are sufficient to accurately conform the radiation dose delivered to target organs on mice. The irradiated field size is affected principally only by the divergence of the beam. In general, the beam has appropriate dosimetric characteristics to accurately deliver dose to organs inside the mice’s bodies.  Using multiple beams delivered from a variety of angular directions, targets as small as 0.1 cm may be irradiated while sparing surrounding tissue.  This microCT/RT system is a feasible tool to irradiate mice using treatment planning and delivery methods analogous to those applied to humans.

I. INTRODUCTION

Small animals are key components of preclinical research. Mice have been employed for many years as models in which cancer biology may be studied and therapies may be developed prior to translation into the clinic. These vertebrate models have served critical roles in improving cancer detection, understanding carcinogenesis and tumor progression, and evaluating therapeutic response. However, while animal models of cancer have advanced tremendously, models of radiotherapy applicable to these subjects have lagged behind their clinical counterparts. In the early 70’s, Hranitzky et al(Hranitzky et al., 1973) designed an irradiator exclusively for rodents. Their system consisted of two Cs-137 sources of 2600 Ci separated by 27 cm that were placed in a cylindrical lead shield. The system achieved dose rates of 1.1 Gy/min at the intersection of the two sources for field sizes as small as 3 cm in diameter. However, the development of new equipment technology and cancer treatment techniques for humans has limited the relevance of early small animals radiation systems. Currently, standard radiation facilities for humans are equipped with rotating linear accelerators and multi-leaf collimators to deliver intensity modulated radiotherapy (IMRT). Radiation systems are also equipped to account for inter- and intra-fraction organ motion. To reduce the uncertainty in the dose delivery, treatments are monitored with new techniques (Dawson and Jaffray, 2007; Jaffray, 2005; Bouchet et al., 2001) such as on-board kV-cone beam computed tomography (CBCT) and optical-guided 3D ultrasound techniques, creating a new treatment modality dubbed image-guided radiotherapy (IGRT). Dose shaping and monitoring technology of this type, while commonplace in the clinic, is not to be found in the laboratory.

In order to rectify this situation and implement clinically-relevant radiotherapy methods for small animal models of disease, treatment facilities that accommodate the size of these small animals and have dosimetric properties similar to those of radiation systems used to treat tumors in humans must be developed.  Several groups are working to integrate radiation treatment planning and 3-D conformal radiation therapy systems with multimodality imaging capabilities exclusively for rodents (Stojadinovic et al., 2007; Stojadinovic et al., 2006b; Chow and Leung, 2007; Deng et al., 2006; Tryggestad et al., 2006; Ford et al., 2006; Stojadinovic et al., 2006a; Jaffray et al., 2006). In addition, other groups are modifying existing radiation systems to conduct partial body irradiation in rats (Khan et al., 2003; Khan et al., 1998; DesRosiers et al., 2003; Sugiyama et al., 2005; Takahashi et al., 2002; Swicord et al., 1999). Recently, a novel microCT-based IGRT system equipped with a variable-aperture collimator has been proposed (Graves et al., 2007a; Zhou et al., 2007). This device is able to perform 3-D conformal therapy in mice with field sizes as small as 1 mm, making it suitable for irradiation of small targets such as prostate, liver, bladder or lungs in mice.  The unit is based on a commercial microCT scanner, retrofitted with a collimation system consisting of two stages each with an adjustable hexagonal iris (see figure 1).

The two irises are offset by 30° so that when the openings of the two irises are properly adjusted to the same size, a dodecagonal beam profile is formed. While still in development, this system has the potential to deliver intensity modulated radiotherapy by varying the field diameter, beam angle and bed position during the course of radiation delivery.

Commissioning a microCT for mouse irradiation is a challenge due to space restrictions, small field sizes, and detector limitations. Thimble ion chambers may be used to measure output but they are limited to field sizes larger than 4x4 cm2. Alternatively, radiochromic film has been used previously for dosimetry of small fields (Stojadinovic et al., 2005; Deng et al., 2006; Stojadinovic et al., 2006b). Because of its high spatial resolution in two dimensions, radiochromic film is the most appropriate detector to evaluate dose distribution delivered by pinpoint beams. In addition, no specialized film processing is needed, the film is insensitive to light, and it can be cut in any size or shape. Radiochromic film also offers such dosimetric advantages as tissue equivalency. Gafchromic EBT film, for instance, has weak energy dependence (Butson et al., 2006; Fuss et al., 2007), low dose rate dependence (Fuss et al., 2007) and does not exhibit temporal growth of optical density (Fuss et al., 2007; Todorovic et al., 2006; Cheung et al., 2005). A combination of measurements taken with ion chamber and gafchromic EBT film may lead to a feasible absolute and relative dosimetry of small animal irradiators. The main goal of the current study was to commission a microCT system equipped with variable-aperture collimator for small animal radiation therapy. Ion chamber and gafchromic film were used to measure the dose rate at the isocenter of the microCT as a function of depth for different field sizes.  In addition, we quantified the flatness and symmetry of the X-ray beam and evaluated the spatial and dosimetric accuracy of the dose delivered.

II.   METHODS AND MATERIALS

II.A  Output of the microCT/RT system

A GE eXplore Locus microCT (GE Healthcare, Milwaukee, WI) has been modified to irradiate mice in style similar to tomotherapy (figure 2a). A collimator with an adjustable variable-aperture has been adapted in order to limit the beam to fields as small as 1 mm in diameter.  The output of the 120 kVp photon beam provided by the GE microCT was measured following the guidelines of the in-air method presented in the AAPM protocol for 40-300 kV x-ray beam dosimetry in radiotherapy and radiobiology, TG-61 (Ma et al., 2001). The protocol is based on ionization chambers calibrated in air in terms of air kerma. The ion chamber must have an air-kerma calibration factor, Nk, for the beam’s quality to be measured. The data was collected using a cylindrical PTW N30006 ion chamber (PTW-Freiburg, Freiburg, Germany), and an Inovision Model 35040 electrometer (CNMC Company Inc. Nashville, TN). The ion chamber was placed at the scanner isocenter, in air, and irradiated with a 120 kVp photon beam in a field diameter of 6, 5 and 4 cm (figure 2b). The field diameter refers to the field width of the dodecagon formed by the two irises collimator apertures. Measurements were performed with the x-ray tube operating at currents of 25, 40 and 50 mA. Ten readings were performed for each field diameter and tube current setup.

II.B  Film dosimetry and calibration curve

The dose rate for radiation fields smaller than 4 cm in diameter should not be measured using ion chamber due to partial volume effect in the chamber (Laub and Wong, 2003).  Because of its energy independence and high sensitivity, gafchromic EBT film is a good detector for two-dimensional dose measurements at lower photon energies (Butson et al., 2006; Chiu-Tsao et al., 2005). Gafchromic EBT film was used to measure the dose rate for field diameters smaller than 4 cm. The optical density of the film was related to the dose delivered using a calibration curve performed in the same microCT scanner. To create the calibration curve, films were placed at the isocenter of the microCT on top of a solid water slab of 6x6x3 cm3 and irradiated at doses of 0, 0.5, 1.0, 1.5, 2.0, 2.5, 3.0, and 3.6 Gy. One millimeter of solid water was placed on top of the film to ensure both electronic equilibrium and good contact between film and solid water. Films were read in an Epson scanner with the capability to read films using transmitted light. Each film image was saved as a 16 bit grayscale TIFF file at a resolution of 508 dpi. Care was taken to analyze films consistently (same batch, same film orientation, same reading position, same post-irradiation reading, etc). Analysis of the film images was performed in MatLab. Films measurements and analysis were made in accordance with the guidelines suggested by AAPM Task Group 55 (Niroomand-Rad et al., 1998).

II.C  Dose rate as a function of field diameter

Relative film dosimetry was used to find the dose rate at field diameters smaller than 4 cm. Pieces of gafchromic EBT film with dimensions 6x6 cm2 were placed at the scanner isocenter on top of a 6x6x3 cm3 solid water phantom and irradiated using field diameters of 5.0, 4.0, 2.0, 1.0, 0.5, and 0.2 cm. Films irradiated with beams of diameters 5 and 4 cm were used as references to find the relative dose rate for the smaller fields. Measurements were performed three times on different days to analyze the overall uncertainty in the measurement (due to beam variation, detector sensitivity, and positioning).  One millimeter of solid water was placed on top of the film to ensure both electronic equilibrium and good contact between film and solid water. The film phantom was secured tightly to ensure good film-solid water contact. We also used a PTW TN31014 ion chamber, with a diameter of 2 mm and sensitive volume of 0.015 cm3, to detect any bore and collimator scatter as a function of field diameter. The pinpoint chamber was placed at the isocenter and measurements were taken in air for field diameters from 1 cm to 6 cm.

II.D  Dose rate as a function of depth

Dose rate as a function of depth was also investigated using gafchromic EBT film. Film pieces of 6x6 cm2 were sandwiched in solid water slabs of 6x6x0.3 cm3 (see figure 2c) to form a stack of 11 pieces of films (6x6x0.0025 cm3), 11 solid water slabs (6x6x0.3 cm3) and 1 solid water slab (6x6x0.1 cm3). The solid water-film phantom was placed with the top at the isocenter and irradiated with a 120 kVp photon beam at tube current of 50 mA for field diameters of 2.0, 1.0, 0.5 and 0.2 cm.  The solid water-film stack was firmly secured to ensure good film-solid water contact.

II.E Flatness, symmetry, FWHM and penumbra

Flatness, symmetry, penumbra, and full-width half-maximum (FWHM) for beam diameters of 2.0, 1.0 and 0.5 cm were also analyzed using gafchromic EBT film. Flatness of the beam was defined as the maximum variation of the dose, normalized to the central axis, over 80% of the cross-sectional area of the beam profile, and symmetry was calculated as the maximum percentage deviation of the left-side dose from the right-side dose of the beam profile at 80% of the FWHM. The FWHM refers to the distance between the points at each side of the beam profile where the dose is 50% of the maximum dose, and penumbra refers to the lateral distance between the 80% and 20% of the maximum dose on one side of the beam profile (parameters are defined as proposed by AAPM protocol TG-45 (Nath et al., 1994). Films were placed at 2.5 cm above the isocenter and irradiated in solid water with 0.1 cm buildup and 3.0 cm backscatter. To analyze the dose profile changes as a function of depth, films were also sandwiched at 1.0 and 2.0 cm depth in the solid water for the field diameter of 1.0 cm.

II.F Irradiation time calculation

To facilitate calculation of irradiation times needed to achieve desired doses in a subject, the dose rate as a function of depth data was converted to tissue maximum ratio (TMR) data. Irradiation time was calculated the following equation:

(1)

where TD is the total dose delivered at the center of the target, DR(fs, 120kVp,50mA) is the dose rate of the beam at the isocenter at 1.0 mm depth in water for a filed diameter fs, γ = 120 kVp and I = 50 mA. kmA is the tube current factor if x-ray operates at tube current other than 50 mA, TMRi(fs,d) is the tissue-maximum ratio at depth d for a field diameter fs, and n is the number of beams. This formula assumes beams with equal field diameters. To test the accuracy of dose delivery in a single fraction, a 3x3 cm2 piece of gafchromic EBT film was sandwiched between two solid water slabs of 3x3x1.5 cm3 and placed at the isocenter with the film parallel to the beam and the gantry rotation plane (see figure 2d). The film was irradiated with 8 beams, 45° apart, with field diameters of 0.2 cm. The tube was operated at 120 kV and 50 mA.  The dose at the isocenter was manually calculated using Equation 1 to be 1.40 Gy and was compared to that measured by the gafchromic EBT film.

III. RESULTS AND DISCUSSION

III.A Output of the microCT/RT system

The dose rate at the isocenter of the microCT/RT system at the surface of water was measured using the in-air method suggested by AAPM protocol TG-61 (Ma et al., 2001). Table 1 shows the dose rate at the isocenter at the surface of water for field diameters of 4.0, 5.0 and 6.0 cm provided by the 120 kVp photon beam with the tube current operating at 25, 40 and 50 mA. The maximum standard deviation of the measurements is 0.3%.

| Tube current | Dose rate (Gy/min)a |
| --- | --- |
| (mA) |  | 6.0 cm | 5.0 cm | 4.0 cm |
| 50 |  | 2.13 | 2.08 | 2.04 |
| 40 |  | 1.72 | 1.68 | 1.62 |
| 25 |  | 1.11 | 1.08 | 1.02 |
| a Dose rate at isocenter at surface of water for a photon beam of 120 kVp |

The dose rate is high enough to deliver therapeutic doses of radiation in reasonable periods of time. However, there are two considerations that should account when irradiating small animals with the system; 1) the duty cycle of the system and 2) generator overheating.  The system provides a pulsed beam and a full cycle of one pulse includes the on-time, the decay time and the camera read-out time, I included because of the design of the system as an imaging device. The on-time is the time when the beam is delivering the dose and is approximately 10% of the full cycle time. Therefore, the irradiation time is actually 10 times higher than the time calculated using dose rate presented in table 1. Secondly, the generator of the system overheats when delivering dose for long periods of time. However, a careful study of the heating behavior has been performed to establish dose delivering patterns without overstressing the generator. In general, including positioning, scanning and dose delivering, the treatment time is approximately 8 min per Gy delivered.

III.B Film dosimetry and calibration curve

Dosimetry for field diameters smaller than 4 cm was performed using gafchromic EBT film. Figure 3 shows the graph that associates the optical density and the dose delivered to the film. The films were read in a 16-bit flat bed scanner using transmitted light. The optical density of the film is defined as the logarithm of the light transmitted through the film. This is

(2)

where the I0 is the light transmitted through an unused film (base + fog only) and I is the light transmitted through the irradiated film. Log10(I) and Log10(I0) are represented by the pixel value of the irradiated film and the unused film respectively.

The curve can be fit well with a polynomial of third order.

III.C Dose rate as a function of field diameter

Figure 4 shows the dose rate as a function of field diameter. The dose rate ranges from 1.56 Gy/min to 2.08 Gy/min for field diameters between 0.1 and 5 cm, respectively. The dose rate decreases about 26% when reducing the collimator aperture from 5.0 cm to 0.1 cm. The dose rate decreases because of the scatter reduction when photons interact with matter. The pinpoint chamber measurements suggest that the contribution of the bore and collimator scatter to the primary beam is negligible. Measurements with field diameters from 1 cm to 5 cm did not show variation. Standard deviation of the measurements in that field diameter range was within 0.9% which can be attributed to the beam fluctuation. AAPM TG-61 protocol presents backscatter factors for different energies, source-surface distances and field diameters. For the same half value layer of the microCT beam and for a source-surface distance equal to the isocenter distance, the backscatter factor for field diameters of 5, 2 and 1 cm are 1.242, 1.12 and 1.059 respectively.   The backscatter ratios of the field diameter of 1 cm with respect to the 2 and 5 cm are 0.945 and 0.853 respectively. The ratios of the measured dose rate of the field diameter of 1 cm with respect to the 2 and 5 cm are 0.945 and 0.832 respectively. This suggests that the variation of the dose rate as a function of field diameter is basically due to scatter when photons interact with the solid water. The protocol does not present any data for field diameters smaller than 1 cm. However, if we assume the field diameter of 0.2 cm as a pencil beam, the backscatter factor is 1 and the ratio with the backscatter factor for a field diameter of 5 cm is 0.80. Using the measured dose rate, that ratio is 0.75. However, at field diameters of 0.2 cm, the dose rate might be also affected for other factors like penumbra and reduction of photons passing through the aperture. In conclusion, even for very small field diameters, the dose rate is still high enough to irradiate targets of 1 mm in diameter. The error bar in plot presented in figure 4 represents the average standard deviation obtained in all points (3.6%).

III.D  Dose rate as a function of depth

Figure 5 shows the dose rate as a function of depth for field diameters of 2.0, 1.0, 0.5 and 0.2 cm. The depth of maximum dose is somewhere between the surface and 1 mm depth. However, the film has neither the spatial resolution nor the sensitivity to measure the depth of maximum dose. One millimiter depth will be taken as the depth of maximum dose based on the following facts; 1) 1 mm depth is the shallowest depth that we can measure using film dosimetry, 2) uncertainty on the detector is higher than any dose variation between the surface and 1 mm depth, and 3) irradiation time for dose delivered in mice will be calculated at a minimum of 1 mm depth even for subcutaneous tumors.  For field diameters of 1.0 cm and 0.5 cm, appropriate field sizes for mice irradiation, the dose rate decreases about 10% per every 5 mm in depth. The dose rate is still high enough to irradiate targets inside of mice’ bodies considering that the average radius of a mouse’s body is about 1.25 cm. The dose rate drops slower for field diameters larger than 1.0 cm and slightly faster for field diameters smaller than 0.5 cm. The depth dose curve for 2.0 cm presents a shoulder which is not evident at a field diameter of 0.2 cm. That is due to the fact that the backscatter for field diameters of 2.0 cm is significant higher than backscatter for 0.2 cm.

III.E Flatness, symmetry, FWHM and penumbra

Figure 6a shows the 120 kVp photon beam dose profile at 1 mm depth for field diameters of 2.0, 1.0 and 0.5 cm. Field diameters were set at the isocenter but the film was placed at 2.5 cm above the isocenter to increase dose rate. Therefore, the actual field diameters are 1.86, 0.93 and 0.46 cm.  Table 2 presents the measured values of penumbra and FWHM. It also provides the calculated values of flatness and symmetry of the beam.

|  | Field diameter |
| --- | --- |
|  | 2.0 cm | 1.0 cm | 0.5 cm |
| Flatness | 3.4% | 3.6% | 3.7% |
| Symmetry | 1.0% | 1.0% | 0.1% |
| Penumbra (Film) (mm) | 0.60 ± 0.05 | 0.45 ± 0.05 | 0.45 ± 0.05 |
| Penumbra (CT detector) (mm) | 0.50 ± 0.05 | 0.50 ± 0.05 | 0.50 ± 0.05 |
| Field diameter at film (mm) | 1.86 | 0.93 | 0.46 |
| FWHM | 1.83 ± 0.05 | 0.92 ± 0.05 | 0.46 ± 0.05 |

The beam presents a good symmetry at field diameters smaller than 2.0 cm in diameter. The heel effect is insignificant at these small field diameters. It starts to be notable at field diameters larger than 4.0 cm, however such field sizes are not relevant for mouse treatments. The dose profiles exhibit shoulders but are flat enough to deliver dose homogenously at 75% of the cross-area.  The beam has a very sharp penumbra. However, dose rate can dramatically drop at field diameters smaller than 0.2 cm. At field diameters as small as 0.2 cm, the dose rate will decrease not only due to the scatter reduction but also due to penumbra effect.  The field diameter set in the collimator and the field diameter measured with film at 1 mm depth agreed within 1.6%. The radiation field diameter increases to 0.95 and 1.01 cm at depths of 1 and 2 cm respectively (see Fig 6b). This increase is mainly due to the divergence of the beam with a little contribution of lateral photon scatter.  The uncertainties in the penumbra and FWHM are the size of the pixel in the images. In general, the photon beam has good dosimetric characteristics to homogenously irradiate targets inside the mice’s bodies.

III.F Dose delivering testing

The film-solid water phantom was irradiated with 8 beams with diameters of 0.2 cm. The calculated dose delivered at the intersection of the beam axes was 1.40 Gy. Using the calibration curve and the film optical density (OD) in the interception of the beam axes as the input, the measured dose delivered to the film was 1.36 Gy. The difference between the calculated and measured dose is 2.8%.  Figure 7 shows the irradiated film and the image of the dose distribution in the film. This single dose testing demonstrated that the dosimetry of the system was consistent and that it can be used to calculate irradiation time when delivering dose to internal organs in mice.

III.G Error analysis

The dose delivered to a specimen is calculated based on the air kerma rate at the isocenter of the microCT. The air kerma rate has been measured as accurately as possible but in general, the result of the measurement is only an approximation or estimation of the true value. The result is complete only when the confidence level of the value is well known. In other words, the result must be accompanied by a quantitative statement of its uncertainty associated with the measurements (Taylor and Kuyatt, 1994). The total uncertainty should include random and systematic errors in the measurement process. The total uncertainty of the air kerma rate at the isocenter of the microCT represents the combined uncertainty of type A and type B uncertainties determined over the entire calibration procedure. It includes uncertainty due to fluctuations in the photon beam, uncertainty of the detector position, and uncertainty of detector sensitivity. The uncertainties involved in the dose rate measurement, using an ion chamber, are shown in table 3. The type A uncertainties were determined based on statistical analysis (standard deviation), whereas the type B uncertainty was based on previous measurement data, manufacturer’s specifications, and data provided in calibration and other reports. Scientific judgment was also used to estimate or assign probability distributions over the limits in parameters provided by manufacturer’s specifications. The values for the standard uncertainty type A and type B are 0.3% and 2.4%, respectively. The value for the combined uncertainty is 2.4%. The expanded uncertainty at a coverage factor of k = 2 is, therefore, 4.8%. When using film dosimetry, the overall uncertainty was estimated in 3.6% (includes uncertainties in taking the measurements and reading the film). The expanded uncertainty at a cover factor of k = 2 is therefore 7.2 %.

|  | Type A (%) | Type B (%) |
| --- | --- | --- |
| Air kerma calibration factor |  | 2.0 |
| x-ray beam fluctuation |  | 0.9 |
| Position |  | 1.0 |
| Ion chamber readings | 0.3 |  |
| Electrometer |  | 0.3 |
| Uncertainty | 0.3 | 2.4 |

IV. CONCLUSION

The main goal of this study was to investigate the dosimetric characteristics of the 120 kVp photon beam of a GE Xplore Locus microCT modified to have therapy capabilities. When the x-ray tube operates at the maximum tube current, 50 mA, the dose rate at the isocenter ranges from 2.13 to 1.56 Gy/min at the surface of water for field diameters between 6.0 and 0.2 cm. This dose rate is high enough to deliver dose in mice in reasonable period of time. For field diameters ranging between 1.0 and 0.5 cm, the dose rate decreases about 10% per every 5 mm depth, which make the beam feasible to irradiate targets inside of mice’s bodies.  The beam has good flatness, symmetry and sharpness to accurately conform dose distribution on target organs on mice and the irradiation field size is affected basically only by the divergence of the beam. The method to calculate irradiation time presented in this paper is mainly based on the dose rate at the central axis of the beam and the dose delivered at the center of the target. Testing was performed only in two dimensions but due to the acceptable flatness and symmetry of the beam, we can suggest that the dose distribution has acceptable homogeneity over 80% of the irradiated volume. Efforts are now in progress to produce a treatment planning system for this microCT/RT device using RT_Image (Graves et al., 2007b), which will incorporate the commissioning data presented here and calculate irradiation time based on the integral dose delivered to the gross tumor volume (GTV). The work presented in this paper is being used as the dosimetric platform to validate the beams generated with Monte Carlo simulation which will be used for calculation in the treatment planning software. We are investigating strategies to increase the dose rate of the system. The dose rate ranges from 2.13 to 1.56 Gy/min but the delivering time is restricted by two factors, the duty cycle of the system and the limitation in heating circulation on the generator housing. Further modifications on the system (cooling, tube capacity, etc) have been proposed to improve the delivering rate. At this time the system is capable of delivering dose at a rate of approximately 1 Gy per every 8 minutes without overstressing the system. This delivery rate includes positioning, scanning, image reconstruction and dose delivering, In general, the beam has appropriate dosimetric characteristics to accurately deliver dose to organs inside mice. The dose is precisely delivered in a tomotherapy style with field diameters as small as 0.1 cm.  The 3-D stage, collimator and gantry are computer-controlled which make the system ready for intensity modulated radiation therapy capabilities.  This microCT/RT system is a feasible tool to irradiate mice with modalities and treatment plans similar to those performed in human.

V. REFERENCES

Bouchet L, Meeks S, Goodchild G, Bova F, Buatti J and Friedman W 2001 Calibration of three-dimensional ultrasound images for image-guided radiaiton therapy Phys. Med. Biol. 46 559-77

Butson M, Cheung T and Yu P 2006 Weak energy dependence of EBT gafchromic film dose response in the 50 kVp-10 MVp X-ray range Applied Radiation and Isotopes 64 60-2

Cheung T, Butson M and Yu P 2005 Post-irradiation colouration of Gafchromic EBT radiochromic film Phys. Med. Biol. 50 N281-5

Chiu-Tsao S-T, Ho R, Shankar R, Wang L and Harringson L 2005 Energy dependence of response of new high sensitivity radiochromic films for megavoltage and kilovoltage radiation energies Med. Phys. 32 3350-4

Chow J and Leung M 2007 Treatment planning for small animal using Monte Carlo simulation Med. Phys. 34 4810-7

Dawson L and Jaffray D 2007 Advance in image-guided radiation therapy J. Clin. Oncol. 25 938-46

Deng H, Kennedy C, Armour E, McNutt T, Tryggestad E, Ford E, Iordachita I, Kazanzides P, Huang J and Wong J 2006 The small-animal radiation research platform (SARRP): Focused pencil beam dosimetry Med. Phys. 33 2241

DesRosiers C, Mendonca M, Tyree C, Moskvin V, Bank M, Massaro L, Bigsby R, Caperall-Grant A, Valluri S, Dynlacht J and Timmerman R 2003 Use of the leksell Gamma Knife for localized small field lens irradiation in rodent Technol. Can. Res. Treat. 2 449-54

Ford E, Kennedy C, McNutt T, Armour E, Iordachita I, Kazanzides P and Wong J 2006 The small animal radiation research platform: Benchtop cone-beam CT Med. Phys. 33 2017

Fuss M, Sturtewagen E, DeWagter C and Georg D 2007 Dosimetric characterization of Gafchromic EBT film and its implication on film dosimetry quality assurance Phys. Med. Biol. 52 4211-25

Graves E, Chatterjee R, Keall P, Gambhir S, Contag C and Boyer A 2007a Design and evaluation of a variable aperture collimator for conformal radiotherapy of small animals using a microCT scanner Med. Phys. 34 4359-67

Graves E, Quon A and Loo B 2007b RT_Image: An Open Source Tool for Investigating PET in Radiation Oncology. Technology in Cancer Research and Therapy 6 111-21

Hranitzky E, Almond P, Suit H and Moore E 1973 A Cesium-137 irradiator for small laboratory animals Radiology 107 641-4

Jaffray D 2005 Emergent technologies for 3D image-guided radiation delivery Semin. Radiat. Oncol. 15

Jaffray D, Moseley D, Chow J, Kim S, Ansell S, Wilson G and Chiarot C 2006 An imaged-guided irradiator for pre-clinical radiation therapy studies Med. Phys. 33 2241

Khan M, Hill R and Van_Dyk J 1998 Partial volume rat lung irradiation: an evaluation of early DNA damage Int. J. Rad. Onc. Biol. Phys. 40 467-76

Khan M, Van_Dyk J, Yeung I and Hill R 2003 Partial volume rat lung irradiation; assesment of early DNA damage in different lung regions and effect of radical scavengers Radiother. Oncol. 66 95-102

Laub W and Wong T 2003 The volume effect of detectors in the dosimetry of small fields used in IMRT Med. Phys. 30 341-7

Ma C, Coffey C, DeWerd L, Liu C, Nath R, Seltzer S and Seuntjens J 2001 AAPM protocol for 40-300 kV x-ray beam dosimetry in radiotherapy and radiobiology Med. Phys. 28 868-93

Nath R, Biggs P, Bova F, Ling C, Purdy J, Geinj J and Weinhous M 1994 AAPM code of practice for radiotherapy accelerators: Report of AAPM Radiation Therapy Task Group No 45 Med. Phys. 21 1093

Niroomand-Rad A, Blackwell C, Coursey B, Gall K, Galvin J, McLaughlin W, Meigooni A, Nath R, Rogers J and Soares C 1998 Radiochromic film dosimetry: Recommendations of AAPM Radiation Therapy Committee Task Group 55 Med. Phys. 25 2093-115

Stojadinovic S, Hope A, Vivic M, Mutic S, Deasy J, Cui J, Khullar D, Parikh P, Esthappan J, Grigsby P and Low D 2006a MicroRT/MicroRTP: A conformal small animal planning and radiation system Med. Phys. 33 2273

Stojadinovic S, Low D, Hope A, Vivic M, Deasy J, Cui J, Khullar D, Parikh P, Malinowski K, Izaguirre E, Mutic S and Grigsby P 2007 MicroRT- Small animal conformal irradiator Med. Phys. 34 4706-15

Stojadinovic S, Low D, Vivic M, Mutic S, Deasy J, Hope A, Parikh P and Grigsby P 2006b Progress toward a microradiation therapy: Small animal conformal irradiator

Med. Phys. 33 3834-45

Stojadinovic S, Low D, Vivic M, Mutic S, Gribsby P, Deasy J and Hope A 2005 Progress towards a MicroRT small animal conformal irradiator Med. Phys. 32 2063

Sugiyama K, Nakatsuji A, Ochiai H, Esumi H and Matsumura Y 2005 Mechanism of impairment in wound healing of electron beam-irradiated rat skin Wound Rep. and Reg. 13 A5-A(1)

Swicord M, Morrissey J, Zakharia D, Ballen M and Balzano Q 1999 Dosimetry in mice exposed to 1.6 GHz microwaves in a carrousel irradiator Bioelectromagnetics 20 42-7

Takahashi S, Sun X Z, Kubota Y, Takai N and Nojima K 2002 Histological and elemental changes in the rat brain after local irradiation with carbon ion beams J. of Rad. Res. 43 143-52

Taylor B and Kuyatt C 1994 Guidelines for evaluating and expressing teh uncertainty of NIST measurement results National Institute of Satndards and Technology

Todorovic, M, Fisher, M, Cremers, F, Thom, E, Schmidt and R 2006 Evaluation of Gafchromic EBT prototype B for external beam dose verification Med. Phys. 33 1321-8

Tryggestad E, Armour E, Deng H, Ford E, Huang J, Iordachita I, Kazanzides P, Kennedy C, McNutt T, Verhaegen F and Wong J 2006 The small-animal radiation research platform (SARRP): Commissioning a 225 kVp "small field" x-ray source for Monte Carlo-based treatment planning Med. Phys. 33 2241

Zhou H, Chatterjee R, Contag C, Gambhir S, Boyer A, Keall P and Graves E 2007 Development of a variable-aperture collimator for small animal radiation therapy Med. Phys. 34 2436

# Text boxes

a)

b)

c)

d)

Figure 1. The variable-aperture collimator in the microCT/RT system uses two discs with hexagonal shape that open and close like iris camera.  The discs are rotated 30º one respect to the other to form a field of dodecagonal shape.

Figure 2.  a) Targets for irradiation are placed at the isocenter of the microCT and treated with beams from different angles.  b) To measure the dose rate, the ion chamber was placed at the isocenter, in air, and irradiated with a photon beam of 120 kVp with a field diameter of 6 cm. c) To measure the percentage depth dose, gafchromic films were sandwiched in 6x6x0.3 cm3 solid water slabs and the top surface of the solid water was placed at the scanner isocenter. d) A single dose fraction was delivered to a gafchromic film sandwiched in solid, centered at the isocenter and placed parallel to the beam and the gantry rotation plane.

Table 1.  Dose rate at the isocenter of the microCT/RT system. The dose rate is high enough to deliver therapeutic doses to mice in reasonable exposure times.

Figure 3. Calibration curve used to associate the dose delivered to the film and its optical density. The curve can be fit to polynomial of third order.  The error bar represents the standard deviation in reading the film and calculating the optical density, STD = 0.0054.

Figure 4. Dose rate as a function of field diameter.  The dose rate decreases as the field diameter decreases due to scatter reduction. The contribution of the scatter to the primary beam is more significant at low energies than at high energies.  The error bar represents the overall uncertainty in the film measurements, 3.6%.

Figure 5.  Dose rate as a function of depth for field diameters of 2, 1, 0.5, and 0.2 cm.  For field diameters of 1.0 cm and 0.5 cm, most commonly field diameter range used in mice irradiation, the dose rate decreases about 10% per every 5 mm in depth. The dose rate is still high enough to irradiate targets inside of mice’ bodies considering that the average radius of a mouse’s body is about 1.25 cm. The error bar represents the overall uncertainty in the film measurements, 3.6%.

Figure 6.  a) Dose profile at 1 mm depth for field diameters of 1.86, 0.93 and 0.46 cm (120 kVp photon beam). b) Dose profile at 0.1, 1 and 2 cm depth for a field diameter of 0.93 cm. The FWHM of the beam increases to 0.95 and 1.01 cm at depths of 1 and 2 cm respectively. Penumbra, flatness and symmetry do not have significant changes.

Table 2. Flatness, symmetry, FWHM and penumbra for filed diameters of 1.86, 0.93 and 0.46 cm. Measurements were performed with gafchromic EBT film at 1 mm depth in solid water. In general, the photon beam has good dosimetric characteristics to homogenously irradiate targets inside the mice’s bodies.

Figure 7.  a) Film irradiated with 8 beams with diameters of 0.2 cm. b) Dose distribution in the film. The difference between the calculated and measured dose is 1.5%. This single dose testing proved that the dosimetry of the system was consistent and that it can be used to deliver dose to internal organs in mice.

a)

b)

a)

b)

Table 3. Uncertainty of the dose rate measured at the isocenter of the microCT using ion chamber. Uncertainties type A are determined based on statistical analysis, whereas uncertainties type B are based on previous measurements, manufacturer’s specifications, and data provided in calibration and other reports. The combined uncertainty and the expanded uncertainty at a cover factor of k=2 are 2.4% and 4.8% respectively.